ANTENNA MECHANICAL MAINTENANCE KPVLBA 6/5-11/03

APEX INSPECION: In very good condition. FRM 2-YEAR INSPECTION PM.---Done, replaced one bellow. Flex shaft not replaced. Install INA zirks.---All are installed.

FRM INA BEARING CHECK 50 LBS PULL ON PRIMARY SIDE

| Primary Travel: +.0015 | Secondary Side: -.0005 |
| --- | --- |

FRM INA BEARING CHECK 50 LBS PULL ON SECONDARY SIDE

| Primary Travel: -.001 | Secondary Side: +.001 |
| --- | --- |

SUBREFLCTOR INSPECTION: Mounting Hardware.---OK. Dipole.---OK. Peeling paint/delaminating.---OK.

QUAD LEGS INPECTION: Guy wires.---OK. Turn buckle.---OK. Cotter pins.---All in place. Quad leg flange bolts.---All tight.

DISH INSPECTION: Anemometer mounts/operation.---Both mounts ok, cups slightly dented. Panel surfaces.---No dents, paint in good condition. Spot check panel bolts.---Most are tight, a few just ¼ turn.

FEED CONE INSPECTION: Feed cone surface/insulation.--- Caulked where needed, and painted with elastomeric. FE feed windows.---All clear. FE feed mounting hardware.---All tight with no rust. FE feed heaters. Good, no rust. Dichroic reflector cracking/delaminating.---No delamintation, 1 blister in center of panel. Ellipsoid dish/mount.---Realigned/ off ~18”, spring replaced. FC hatch latches replace as needed.--OK.

# ANTENNA MECHANICAL MAINTENANCE KPVLBA 6/5-11/03

| ELEVATION BEARING/AXLE INSPECTION: |
| --- |
| EL. Grease inspection/Encoder side | Looks good, no metal flakes. |
| EL. Grease inspection/DGB side | No metal found in grease. |
| Elevation bearing seals.---OK. Grease elevation bearings.---Done. Check elevation axle for cracks.---No cracks. Listen for popping or grinding.---No popping noticed. |

ELEVATION BULL/PINION GEAR INSPECTION: Bull gear.---Good, well lubricated. Bull gear welds/hardware.--No cracked welds. Pinion gears.---Well lubricated. Lube bull/pinion gears as needed.---Done. Check stow pin.---Checked, needed safety pin.

ELEVATION MOTOR AND GEARBOX INSPECTION: EL.#1 seals/coupling/internal gears.---Seals & couplings ok, gears not checked. Spider replaced. EL.#2 seals/coupling/internal gears.---Replaced spider, gears in good shape ~.002 wear. Gearbox heaters/lube pumps.---All working. Flow gauges/fish eyes.---Cleaned fish eyes. Change gearbox oil as needed.---Oil changed on both gearboxes. Listen for popping or grinding.---No popping noticed.

ELEVATION PLATFORM INSPECTION: Swinging catwalk.---Looks good. Hoist.---Good. Install VTX condenser toe guard.---Done.

ANTENNA STRUCTURAL INSPECTION: Spot-check structural bolts.--Tightened EL pillow block bolt, (encoder side). VTX walk way.---Good. Antenna structural welds.---Great, no cracks. Antenna backup/lower structure.---Good.

ANTENNA MECHANICAL MAINTENANCE KPVLBA 6/5-11/03

AZIMUTH BEARING/GREASE INSPECTION

|  | INNER | OUTER |
| --- | --- | --- |
| D1 | Slight metallic, very few small metal flakes. | Good. |
| D2 | Little flakey, slight scratches on races. | No signs of metal, very smooth races. |
| I1 | Good, very small linear scratches on races. | Very small scratches on races. |
| I2 | Changed bearing, LA used bearing. | Good, few metal flakes, small scratches. |
| Bearing races.---Very slight linear scratches on all races. Close pillow blocks & grease.---Done. |

| AZIMUTH WHEEL INSPECTIONS: Wheel to structure clearances.---Good clearance. Wheel/Axle bolt tightness.---OK. Listen for popping or grinding.---No popping noticed. |
| --- |
| AZIMUTH DRIVE WHEEL ALIGNMENT |
|  | D1 | D2 |
| Horizontal error | 0d 00’18” | 0d 00’57” |
| Vertical error | 0d01’02” | 0d01’47” |
| Radius error | .317” | .135” |

AZIMUTH MOTOR AND GEARBOX INSPECTION: AZ#1 seals/coupling/internal gears.---Gears in good condition, .002 wear. Spider replaced. AZ#2 seals/coupling/internal gears.---Gears not checked, spider will be replaced w/motor c/o. Gearbox heaters/lube pumps.---OK. Flow gauges/fish eyes.---OK, fish eye’s cleaned. Change gearbox oil as needed.---Both gearboxes oil was changed. Listen for popping or grinding.---No popping noticed.

PINTLE BEARING INSPECTION: Cable wrap & cables.---Great shape. Check pocket level.---Done on previous trip. Check for loose bolts.---All tight. Check bearing grease/grease as needed.---Grease OK.

ANTENNA MECHANICAL MAINTENANCE KPVLBA 6/5-11/03

AZIMUTH RAIL INSPECTION: Rail level measurements.---Done. Excessive rail movement.---Minimal. Joint bars and clips.---All in good shape. Antenna foundation/grout/vulchem.---Very good. Re-torque fishplates as needed.---Not needed.

ANTENNA PAINT/INSULATION INPECTION: Antenna paint condition.---Excellent condition. Repair antenna insulation as needed.---Legs look good.

MECHANICAL MAINTENANCE ITEMS TO BE COMPLETED: 1.Replace or repair VTX door.---Patched up, will need replaced. 2.Install QUAD LEG latter and Fall Arrest.---Done. Site techs trained. 3.Realign or replace dichroic panel.---Done, panel was realigned, (~18 inches off).

ANTENNA MECHANICS NOTES:

FRM was well lubricated. We did not pump as much grease into it as we normally do.

All wheel bearings were well lubricated.

ANTENNA SERVO MAINTENANCE KPVLBA 6/5-11/03

ANTENNA SAFETY TESTS: Tests ok.---Done.

| BRAKE HOLDING-TORQUE TESTS ( FT-LBS) |
| --- |
| AZ #1 | 68 ft/lbs |
| AZ #2 | 66 ft/lbs, motor to be replaced. |
| EL #1 | 82ft/lbs |
| EL #2 | 100ft/lbs |

MOTOR INSPECTIONS: Elevation motor/tach couplings.---OK, both motors spider replaced. Elevation brush/commutator.---OK, brushes replaced recently by site techs. Azimuth motor/tach couplings.---AZ#2 grease bleeding into brake, motor to be replaced. Azimuth brush/commutator.---Replaced AZ #1 & AZ #2 brushes.

ANTENNA GROUNDING INSPECTION: Elevation ground cable.---OK. Elevation platform to pintle turret.---OK. Pedestal room.---OK. Azimuth wheel contacts.---OK. Pintle room/cable wrap.---OK.

SERVO RECORDINGS: Elevation response test.---Done. Azimuth response test.---Done. Elevation position loop test.---Done. Azimuth position loop test.--- Done.

ACU AUTO MODE TESTS: System/axis faults.---All good. AZ/EL motor fault status.---All good. Stow commands.---All good. Sychro feedback.---Good. ACU VME communication (comm. dead).-- OK. Restore ACU parameters.---Done.

ANTENNA SERVO MAINTENANCE KPVLBA 6/5-11/03

SERVO MAINTENANCE ITEMS TO BE COMPLETED: 1.Check all motors for ss j-boxes. ---All OK. 2.Replace AZ motor tachs. ---Replaced AZ #2, (10vp-p noise) #1 OK. 3.Install ACU power supply mod.---Done.

SERVO TECHNICIAN NOTES:

Elevation gearboxes heaters were wired in series.

Replaced AZ #2 blower motor.

Replaced EL #1 & #2 motor coupling spiders.

Replaced AZ #1 motor coupling spider.

AZ #2 motor needs replaced.

ANTENNA ELECTRONIC MAINTENANCE KPVLBA 6/5-11/03

APEX/FRM INSPECTION: FRM cable/j-box condition.---Clean, solder joints not shrink wrapped. FRM motor/encoder.---Mod installed. 90/50cm FE/dipole/cables.---Tight, clean and well hung. Replace tie-wraps w/metal as needed. ---Replaced 2.

FEEDCONE IN/OUT INPECTION: Ellipsoid operation.---Adjusted by ant mechs, Operates well. Feed heater operation.---Works. All FE cables.---Neat, tight & orderly. CRYO systems.---With no leaks, replaced 7mm solenoid seal. FE dry air.---Desc. Filters are blue. Replace tie-wrap as needed. Replaced a few.

VTX ROOM IN/OUT INSPECTION: IF/PWR cable/connector bulkhead.---OK, tight. Power bulkhead/rfi filters.---Clean, w/3yer old surface rust, looks good. A/B racks.---Neat, tight, and cables tied.

ANTENNA CABLE INSPECTION: Elevation cable wraps.---Replaced broken ones. All cable trays.---Tight. Elevation platform/limit switches.---Action good. Replace tie-wraps with metal as needed.---Replaced 5 with metal.

PINTLE BEARING INPECTION: Cable wrap/bulkhead.---No evidence of leakage. Limit switches.---Action good. Cable strain relief.---Repaired Teflon on 1 cable.

PEDESTAL ROOM INSPECTION: ACU/FRM logic on UPS? FRM on UPS, see notes*. UPS operation/by-pass switch.---Yes, power outage 6/9 for ~ 10 min. Breakers labeled correctly.---Yes. Dry air compressor.---18/20% run time, good.

ANTENNA ELECTRONIC MAINTENANCE KPVLBA 6/5-11/03

STATION BUILDING INPECTION: Room 100.---Done. UPS operation.---OK, Electrical panels.---Labeled, clean. Onan x/switch.--- Kicked in fine. Room 102.---Looks good, clean. Contempo 1&2 operation.---Works well. Room 103.---Looks good, clean. VME/comm. Rack.---Labels OK. Recorders 1&2.---Good, replaced supply reel hub latch. Fire alarm operation.---OK, checked out. Chatter box & phone operation.---OK. Under floor cables/bulk head.---Floor, clean dry. Room 104.---Looks good/clean. Maser/C&D racks.---Cables tight, racks clean. Under floor cables/bulk head.---Clean, dry, cables /baulk head tight/ good.

WEATHER STATION INSPECTION: Check operation/test alarms.---PM completed. Cabinet/tower.---Secure, no holes, little rust, electronics look fine.

VLBA SITE/STATION GROUNDS INSPECTION: Test generator.---Test fine. Contempo condensers.---Up & running. Station building.---Clean and orderly. Fence/gate/signs/grounds/storage.---Neat, clean and orderly.

ELECTRONIC MANTENANCE ITEMS TO BE COMPLETED: Install FRM focus motor encoder/coupling.---Installed in November of 2002. Install ACU power supply mod.---Completed, & works good.

ANTENNA ELECTRONIC MAINTENANCE KPVLBA 6/5-11/03

FINAL ANTENNA/SITE INSPECTIONS: Spot-check critical PM’s.---Weather station pm completed. Any TT oversights.---None. Site cleanup.---Done Review TT list/problems with site techs.---Done. VLBA operations startup tests.---Looks, good.

VLBA ELECTRONICS TECHNICIAN NOTES:

Ped room PVC feed through broken.

Large Power J-Box in Pintle bearing pit had a lot of corrosion on ground bus.

* AC transformer relay control cable needs replaced.