## Slide 1: ngVLA Pathfinder system Thermal Tests
Hamdi Mani / Sandy Weinreb
06/20/2018
Tests with dummy feed and Stirling cooler
Tests with 1.2-4.2 GHz feed and small (1.5 kW) Sumitomo cooler

## Slide 2
Sunpower GT Stirling Cycle Cryogenic Cooler

## Slide 3
Physical Properties for Dummy_Feed
General Properties:
Material: {Copper}
Density: 8.940 g/cm^3
Mass: 4.610 lbmass
Area: 800.460 in^2 ~ 0.5m^2
Volume: 14.275 in^3

## Slide 5
2 Heat Straps:
1in x 1/8 in x 5in
One strap made of braided copper
The other strap made of solid OFHC Copper

## Slide 8
Configuration # 01

## Slide 9
Configuration # 02

## Slide 10
Configuration # 03

## Slide 11
Configuration # 04

## Slide 12
Configuration # 05

## Slide 13
Configuration # 06

## Slide 15
Feed covered with MLI on sides

## Slide 18
Heat strap connecting the feed to
the Sumitomo 1st stage
Heat strap dimension: 1/8” x 2 “ x 21.5”
OFHC Copper
Covered with 10 layers of MLI

## Slide 19
Temp Sensor mounted on top plate of feed
Temp sensor mounted on bottom plate
Of feed

## Slide 20
Teflon blanket made of 21 layers

## Slide 21
14 mil thick Mylar
Window
Temperatures after 24 Hour cooldown
Top plate of feed: 167K
Bottom plate of feed: 158K
Cooler 1st Stage: 88K
Cooler 2nd stage: 6K
More work is needed to make a better heat strap:
Low thermal resistance heat strap: use OFHC copper material and with low contact resistance

## Slide 22
Low thermal resistance heat strap: use OFHC copper material and with low contact resistance